AGICO Print No. 33

CUREVAL (ver. 5.0)

Purpose

This program serves for off-line processing and graphical representation of the data obtained by measurement of temperature variation of magnetic susceptibility of rocks by means of the CS-3 Apparatus (from room temperature to 700oC) and/or CS-L Apparatus (from -196oC to room temperature) and the KLY-3S/KLY-3 Kappabridge (or CS-23 or CS-2 or CS-L Apparatus and the KLY-2 Kappabridge).

The CUREVAL program was upgraded with the intention to simplify the processing and to provide our customers with more complete data file enabling the data to be plotted in high quality using professional graphic programs. In the data file used in the CUREVAL ver. 3.0 program only time, temperature, measured total susceptibility, and total susceptibility corrected for free furnace/cryostat were stored and the other parameters (such as bulk and normed ones as well as separated ferromagnetic susceptibility component) must have been calculated during processing and could not be simply plotted using professional graphic programs. In the data file used in the CUREVAL ver. 5.0 program all these parameters are also stored and can be easily read by professional graphic programs. The CUREVAL ver. 5.0 program is able to read the data file made by the old program. After the first reading, the datafile is automatically overwritten into the format of the new program.

General

If strongly magnetic rocks are measured whose magnetic signal is much stronger than that of the furnace/cryostat, there are no problems with the data processing and the measured susceptibility values can be used directly in the interpretation. However, if the signal from rock specimen is comparable to, or even weaker than, that of the furnace/cryostat, problems arise from the interference of these signals.

The following analysis will be made in terms of the total susceptibility introduced by Jelínek (1977) which is more convenient than bulk susceptibility, because the former provides us with an information of the magnetic signal regardless of the actual volume of the measured specimen. The total susceptibility is defined as follows

kT = (V/Vo) kB ,

where kT is the total susceptibility, kB is the bulk susceptibility, V is the actual volume of the specimen and Vo is the nominal volume of the specimen (in the case of KLY-2 or KLY-3S/KLY-3 Kappabridge with standard pick up unit Vo = 10 cm3).

In order to avoid these interferences we have developed the following technique (Hrouda 1994). Prior to measuring rock specimens, the temperature variation in total susceptibility of the free furnace/cryostat is investigated. Then, the curve of this variation is smoothed using the method of running means (from 5 points) and the smoothed data is stored in the file. Afterthat, the thermal variation in total susceptibility of the furnace/cryostat with the specimen is measured. Through linear interpolation between neighbouring temperatures on the smoothed free furnace/cryostat curve the total susceptibility of the free furnace/cryostat is calculated for each the temperature in which the susceptibility of the specimen is determined and the free furnace/cryostat total susceptibility is subtracted from the total susceptibility of the furnace/cryostat with specimen. Then, knowing either the volume or mass and density of the specimen, one can calculate the variation in bulk susceptibility of the specimen.

In weakly magnetic rocks whose paramagnetic fraction contributes substantially to the rock susceptibility the thermomagnetic curve is a composite curve arising from the superposition of the paramagnetic hyperbola and the ferromagnetic straight line of zero or very low slope. Hrouda (1994) uses the above susceptibility behaviour for separation of the paramagnetic and ferromagnetic susceptibility components. He supposes that the ferromagnetic susceptibility is more or less constant in the temperature interval between room temperature and 200oC; through fitting the hyperbola to a thermomagnetic curve in the above temperature interval using the least squares method one can resolve the susceptibility into components and determine errors in this resolution. Hrouda, Jelínek & Zapletal (1997) extended this method to the case when the ferromagnetic susceptibility is represented by straight line of non-zero slope.

File Structure

The data obtained by measurement are stored in random access ASCII files made automatically during measurement. Each file contains the data of one specimen and has the extension CUR for high temperature data and CLW for low temperature data. In this file, the first record contains the abbreviations of the parameters stored in individual columns. The second and other records contain the values of the measured and calculated parameters specified in the following table

Abbreviation 	Length 	Parameter stored

TEMP 		9 bytes 	temperature (in degrees centigrade)

TSUSC 	9 		uncorrected total susceptibility

CSUSC 	10 		total susc corrected for empty furnace

NSUSC 	6 		susceptibility normalized by maximum value

BULKS 	12 		bulk susceptibility

FERRT 	9 		total susc of separated ferromagnetic comp.

FERRB 	9 		bulk susc of separated ferromagnetic comp.

TIME 		7 		time of measurement (in seconds) (AUXI) 9 auxiliary data to be 					specified below

(EOL) 		2 		end of line mark

The column containing the auxiliary data is not headed AUXI, but by the name of the free furnace for which the measured data are corrected (F20056 in our example shown below). The suceptibilities except the normalized ones are given in the order of 10-6 [SI]. The file appears as in the following example

TEMP	   TSUSC    CSUSC    NSUSC    BULKS    FERRT    FERRB    TIME    F20056

21.2 	     -164.4       12.94 	      .650        815.0           2.3       147.2            0          815

21.5          -164.4       12.74          .640        802.4           2.1       135.3          16        .170

23.7          -164.6       12.74          .640        802.4           2.2       140.3          34        .500

26.0          -164.8       12.54          .630        789.8           2.1       132.8          50         3.00

29.2          -164.8       12.61          .634        794.2           2.2       144.1          67              5

32.1          -165.0       12.35          .621        777.8           2.1       133.9          83              2

35.3          -165.0       12.42          .624        782.2           2.3       145.0         101             1

37.9          -165.1       12.27          .617        772.8           2.2       140.9         117           25

The column with auxiliary data contains the following data

Line 	Parameter

1 	name of free furnace for which measured data are corrected

2 	bulk susceptibility of neigbouring specimen of standard  volume measured at room 	temperature (in the order of 10-6 SI)

3 	specimen volume (in cm3)

4 	specimen mass (in grams)

5 	specimen density (in g/cm3)

6 	free furnace subtraction mode

7 	mode of susceptibility normalization

8 	mode of calculation of bulk susceptibility

9 	lower temperature (deg C) for susceptibility resolution

10 	upper temperature (deg C) for susceptibility resolution

11 	total resolved susceptibility using ferro const

12 	ferromagnetic component susceptibility

13 	ferromagnetic component determination error

14 	paramagnetic component susceptibility

15 	paramagnetic component determination error

16 	total resolved susceptibility using ferro sloped

17 	ferromagnetic component susceptibility

18 	ferromagnetic component determination error

19 	paramagnetic component susceptibility

20 	paramagnetic component determination error

Computer Requirements

This program can be run on the IBM-PC type computer equipped with DOS-4.0 and higher, VGA graphic card and having about 1MB free space on the hard disk (on the drive C).

Operation

After starting the program the following offer shows:

MAIN MENU

[1] VIEWING, CORRECTING, AND PROCESSING DATA

[2] QUICK PLOTTING A THERMOMAGNETIC CURVE

[3] COMPLEX PLOTTING OF THERMOMAGNETIC CURVES

[4] TERMINATION OF PROGRAM

[1] VIEWING, CORRECTING, AND PROCESSING DATA

In this case, the file to be corrected must be open (see Appendix 1). Then, the status of the specimen appears on the screen. It looks in the following way in the case of specimen just measured and not processed.

STATE OF SPECIMEN: SpecimenName

Auxiliary (volume) data 			Yes 	Bulk s.

Correction for Free Furnace/Cryostat 	No

Norming of Susc by its max Value 		No

Calculation of Bulk Susceptibility 		No

Ferro/Para Separation 			No

Then, the following offer appears

MAKE THE FOLLOWING OPERATION

[1] Viewing and Correcting Primary Data

[2] Correction for Free Furnace/Cryostat

[3] Susceptibility Norming by Maximum Value

[4] Calculation of Bulk Susceptibility

[5] Ferro/Para Separation

[6] Corrections and Calculations [2] to [5]

[7] Susceptibility Resolution Outprint

[8] Next Specimen

[9] Return to MAIN MENU

[1] Viewing and Correcting Primary Data

This procedure serves for displaying the data in the file and for correcting them. The volume data can be corrected, the measurement data can only be deleted (the whole record for one temperature). First, the offer appears

AUXILIARY SPECIMEN DATA:

BULK 	SUSC 		VOLUME 		MASS 		DENSITY

717 			.167 			.502 		3 			example data

CORRECT

[1] Room Temperature Bulk Susceptibility

[2] Specimen Volume

[3] Specimen Mass and Density

[4] No Correction

After selection, the correct data should be input. This can be repeated until selection [4] is chosen. Then, the data of individual measurements appear on the screen. Using arrows or PageUp and PageDown keys, one can go through the file. After pressing D, No. of record to be deleted is input and this record is deleted from the file. Pressing E terminates the procedure.

[2] Correction for Free Furnace/Cryostat

In this case, the data obtained through measurement of the furnace/cryostat with specimen are processed through subtracting the total susceptibility of the free furnace/cryostat from the total susceptibility of the furnace/cryostat with specimen.

After starting this procedure the following offer shows:

SUBTRACT

[1] Constant Free Furnace/Cryostat Value

[2] Complex Free Furnace/Cryostat Susceptibility

[3] No Correction

The selection [1] is recommended to be used in strongly magnetic specimens. The procedure consists of subtraction of one total susceptibility value of free furnace/cryostat from the measured data. Its advantage is simplicity and unnecessity of the knowledge of the whole thermomagnetic curve of free furnace/ cryostat. First, the file with the specimen data to be corrected is open (see Appendix 1) and then the TOTAL SUSCEPTIBILITY OF FREE FURNACE (in [E-6] SI) should be input. (ONE is written instead of the name of the free furnace/cryostat into the header of the last column of the file.) In the STATE OF SPECIMEN the following information appears:

Correction for Free Furnace/Cryostat 	Yes 	constant subtracted

The processing according to the selection [2] is very accurate. For each temperature of the measured data the free furnace/cryostat susceptibility is calculated from neighbouring points on the smoothed free furnace/cryostat thermomagnetic curve through linear interpolation and this is subtracted from the measured data. First, the file with the specimen data to be corrected is open (see Appendix 1). Then, the following information appears in the uppermost line of the screen:

Free Furnace: FurnaceName Confirm or Select Another.

The free furnace/cryostat name is that is written in the specimen data file. One can use this file for the correction or select another free furnace/cryostat data file. Then, the program proceeds automatically. In the STATE OF SPECIMEN the following information appears:

Correction for Free Furnace/Cryostat 	Yes 	whole furnace subtracted

The technique of the correction (as the number of the above procedure) is written into the line 6 of the last column and the name of the free furnace/cryostat file used for the correction is written as the header of the last column of the specimen data file. It should be noted here that the number 1 indicates the subtraction of constant free furnace/cryostat value, while the complex furnace/cryostat correction is indicated by the number 5, because of the compatibility with the CUREVAL (ver. 2.0) program which worked with more correction techniques, but these were mostly unsused by our costumers.

[3] Susceptibility Norming by Maximum Value

The maximum susceptibility value corrected for free furnace/cryostat is found first. Then, all these susceptibility values are normed by the above maximum susceptibility value and stored in the fourth column of the file. In the STATE OF SPECIMEN the following information appears:

Norming of Susc by its max Value 	Yes 	corrected susc normed

In the case that the susceptibility data corrected for the free furnace/cryostat do not exist, the norming is made of the uncorrected total susceptibility data (and information of that appears on the screen). The results are also stored in the fourth column of the file. In the STATE OF SPECIMEN the following information appears:

Norming of Susc by its max Value 	Yes 	uncorrected susc normed

[4] Calculation of Bulk Susceptibility

The program checks the volume data of the specimen. If there are none, it shows the information INSUFFICIENT DATA, CALCULATION SUPRESSED and terminates the procedured without calculation. If the volume data exist, it gives the following offer

Calculate Bulk Susc From:

[1] Bulk Susc of Neighbouring Specimen at Room Temperature

[2] Total Susceptibility and Specimen Volume

[3] Total Susceptibility and Specimen Mass + Density

or reduced offer according to the existence of the volume data in the file. After selecting the way of bulk susceptibility calculation, the bulk susceptibility is calculated and the results stored in the fifth column of the data file. In the STATE OF SPECIMEN the following information appears accordingly:

Calculation of Bulk Susceptibility 	Yes 	from specimen room susc

Calculation of Bulk Susceptibility 	Yes 	from specimen volume

Calculation of Bulk Susceptibility 	Yes 	from specimen mass+density

[5] Ferro/Para Separation

The heating thermomagnetic curve normed by its maximum value is displayed and the Lower, Upper Temperatures for Hyperbola Fitting are input. Then, the susceptibility resolution into ferromagnetic and paramagnetic component is made using the constant ferro susceptibility method by Hrouda (1994) and the sloped ferro susceptibility method by Hrouda et al. (1997). The paramagnetic susceptibility calculated by the former method is subtracted from the specimen susceptibility and the separated ferromagnetic susceptibility is stored into the file as the total susceptibility (6-th column) and bulk susceptibility (7-th column). The fitted hyperbola is drawn so that one can find the best resolution interval through repeated trials. The results also appear on the screen in terms of total susceptibilities (in the order of 10-6). Kt is rock susceptibility, Kp and Kf are paramagnetic and ferromagnetic components, respectively, and Sp and Sf are errors in the determination of the paramagnetic and ferromagnetic components, respectively. The results of the resolution are also written into the last column of the data file.

If one, after displaying the theromagnetic curve, sees no hyperbola and wishes to avoid the component separation, inputs 0,0 for Lower, Upper Temperatures for Hyperbola Fitting; then the separation is not made.

[6] Corrections and Calculations [2] to [5]

This operation is made, if the corrections and calculations [2] to [5] should be made at once in order to save time and energy.

[7] Susceptibility Resolution Outprint

The data of the susceptibility resolution into ferromagnetic and paramagnetic components can be displayed on screen, printed on paper or written on disk (extension *.CUC) in the form shown by the following example (total and bulk susceptibilities are given in the order of 10-6 [SI]).

Specimen : n20e1_2

Susceptibility Resolution into Components

ferro const 	Total 		Bulk 		Norm

rock susc 	12.73 		801.77 	1.000

ferro comp 	2.36 		148.64 	0.185

ferro error 	0.04 		2.52 		0.003

para comp 	10.37 		653.13 	0.815

para error 	0.06 		3.78 		0.005

ferro sloped

rock susc 	12.88 		811.22 	1.000

ferro comp 	0.58 		36.53 		0.045

ferro error 	0.31 		19.52 		0.024

para comp 	12.29 		774.06 	0.954

para error 	0.32 		20.15 		0.025

[8] Next Specimen

After this selection, new specimen is open in order to make similar corrections and calculations.

[9] Return to MAIN MENU

This selection is made, if one wishes to finish corrections and calculations. The program goes back to the MAIN MENU.

Processing of Free Furnace/Cryostat Data File

The program recognizes whether the file under consideration contains the measurement of specimen or those of the free furnace/cryostat (EMPTY is written in the header of the last column). In the latter case the STATUS has the following form

Free Furnace Data Smoothed: 	Yes (No)

MAKE THE FOLLOWING OPERATION

[1] Viewing and Correcting Primary Data

[2] Smoothing Free Furnace Data

[3] No Correction or Calculation

One may look on the measured data using procedure [1] in the same way as in the case of specimen data or smooth the measured data using the technique of running means (from 5 points). The procedure is made automatically. Then, both the non-smoothed (in red for heating curve and in blue for cooling curve) and smoothed (in white) curves of the free furnace/cryostat appear on the screen normalized by maximum value. The parameter SE which appears in the upper right corner is the mean deviation of the non-smoothed and smoothed values in terms of total susceptibility in the order of 10-6 SI (for details see Hrouda 1994). Pressing ENTER goes back to the menu.

Before smoothing, the measured data are tested through calculating successive differences of measured total susceptibilities. If the differences are larger than 0.5 x 10-6 [SI], they are displayed and one decides whether to accept them or whether to delete the deviated measurements. In the deletion procedure, the succesive numbers of the deviated measurements are highlighted for better orientation.

[2] QUICK PLOTTING A THERMOMAGNETIC CURVE

One only selects the specimen to be plotted and the plotting is made automatically. Plotted are both heating and cooling thermomagnetic curves. Preferrably, the corrected data are used, if they do not exist in the file, the uncorrected data are used. If the free furnace curves are plotted, the uncorrected data are used. The procedure is terminated by pressing ENTER.

If  both high tempreature curves exist, i.e. the heating and cooling ones, two so called alteration factors  are calculated. The are defined as follows

A(40) = 100 (kTc,40 - kTh,40)/kTh,40

A(max) = 100 (kTc,max - kTh,max)/kTh,40

where kTc,40 and kTh,40   are cooling and heating total susceptibilities at 40oC, while kTc,max and kTh,max are are cooling and heating total susceptibilities at the temperature where the factor is maximum.

These factors are positive, if the cooling susceptibility is higher than the heating one and and negative, if the heating susceptibility is higher than the cooling one.

[3] COMPLEX PLOTTING OF THERMOMAGNETIC CURVES

In this case, the thermomagnetic curves are represented graphically on the screen of the computer from where they can be copied either on the paper or into a file (using convenient programs which are not part of the delivery of this program). It is possible to plot more curves into one figure and combine the high temperature curves with the low temperature ones. First, the files to be plotted must be open (see Appendix 1). Then, the following offer appears divided into three blocks

PLOTTING OPTIONS

UNCORRECTED TOTAL SUSCEPTIBILITY

CORRECTED TOTAL SUSCEPTIBILITY

NORMALIZED SUSCEPTIBILITY

BULK SUSCEPTIBILITY

SEPARATED FERRO TOTAL SUSCEPTIBILITY

SEPARATED FERRO BULK SUSCEPTIBILITY

PLOT ANOTHER SPECIMEN(S)

TERMINATION OF PLOTTING

PLOT HEATING CURVE

PLOT COOLING CURVE

PLOT BOTH THE CURVES

PLOT COLOURED CURVES

PLOT WHITE CURVES

The procedure UNCORRECTED TOTAL SUSCEPTIBILITY plots the thermomagnetic curves as they were measured, without any correction. In majority of specimens (weakly to mediately magnetic) it means that the plotted susceptibilitites are negative (due to the diamagnetism of the free furnace/cryostat).

The procedure CORRECTED TOTAL SUSCEPTIBILITY plots the thermomagnetic curves corrected for the free furnace/cryostat (subtraction of the smoothed free furnace/cryostat curve from the measured specimen curve).

The procedure NORMALIZED SUSCEPTIBILITY plots the thermomagnetic curves normalized by the maximum value as they are written in the column 4 of the data file.

The procedure BULK SUSCEPTIBILITY plots the thermomagnetic curves of the bulk susceptibilities as they are written in the column 5 of the data file.

The procedure SEPARATED FERRO TOTAL SUSCEPTIBILITY plots the thermomagnetic curves of the ferromagnetic susceptibility component after subtraction of the paramagnetic hyperbola from sample's thermomagnetic curves (column 6).

The procedure SEPARATED FERRO BULK SUSCEPTIBILITY plots the thermomagnetic curves of the ferromagnetic susceptibility component after subtraction of the paramagnetic hyperbola from sample's thermomagnetic curves and corrected for specimen volume (column 7).

If the selection PLOT COLOURED CURVES is made, the heating curve is plotted in red, while the cooling curve in blue. If the procedure PLOT WHITE CURVES is made, the heating curve is plotted as full white line, while the cooling curve is plotted as dotted white line (this plotting is useful in preparation of hardcopies for print).

After the figure appears on the screen, it can be copied on the paper or written into the file using standard programs which are not supplied with the CUREVAL program. After pressing ENTER the figure disappears and one can choose another type of plot or finish the plotting. In the latter case the starting offer shows on the screen and one can either process another specimen or finish the work with the CUREVAL program.

Note: If the low temperature and high temperature total susceptibility thermomagnetic curves of a specimen are plotted into one figure and the volumes of the respective sub-specimens are different, the latter curve does not appear as a continuation of the former curve, because of the different total susceptibilities measured. If one wishes to plot the low temperature and high temperature curves as one quasi-continuous curve, one has to work with bulk susceptibilities.

References

Hrouda, F., 1994. A technique for the measurement of thermal  changes of magnetic susceptibility of weakly magnetic rocks  by the CS-2 apparatus and KLY-2 Kappabridge. Geophys. J.  Int., 118, 604-612.

Hrouda, F., Jelínek, V. & Zapletal, K., 1997. Refinement of the  technique for susceptibility resolution into ferromagnetic  and paramagnetic components. Geophys. J. Int., 129, 715-719.

Jelínek, V., 1977. The statistical theory of measuring anisotropy  of magnetic susceptibility of rocks and its application.  Geofyzika Brno, 77 pp.

Appendix 1. Opening Data Files

The file opening starts with displaying all the files and subdirectories (<DIR>) in the subdirectory used the last time. The file worked with the last time is highlighted. The desired file or subdirectory is activated through highlighting it using arrows (the last item of the dispayed file list can also be reached through pressing END key). The highlighted file is open by pressing ENTER. If one wishes to change the drive, one highlights the CH DRIVE, presses ENTER and inputs the name of the desired drive. If one highlights the RETURN and presses ENTER the procedure under operation is terminated except the COMPLEX PLOTTING OF THERMOMAGNETIC CURVES procedure.

In the procedure COMPLEX PLOTTING OF THERMOMAGNETIC CURVES, one can plot several curves into one figure. In this case, one has to open several files. This is made by repeated procedure described above. Selecting is terminated by highlighting RETURN and pressing ENTER.

Before working with the data file, the test is made whether the disk containg the above file is not write-protect. If the disk is write-protect, it is indicated by the program and one has to unlock it.

Thermomagnetic curves of a specimen with well defined ferromagnetic and paramagnetic components (left) and of a separated ferromagnetic component (right)

Low temperature and high temperature thermomagnetic curves of hematite plotted in one figure; note the differences between total susceptibilities (left) and bulk susceptibilities (right)